## Sheet (Back)
Root Cause | The first root cause identified relates to a lack of visual management in the raw material storage lcoation, leading to an incorrect part number being picked for packaging. The second root cause identified relates to a lack of visual references to the appearance of the cassette, where the operator is unable to visually verify the cassette's appearance
Action | Description | Due Date | Owner
1 | Update the layout of the propack area of the warehouse to include visual management in the form of visual kanban | 2024-12-20 | Julie Bartoli
2 | Implement a visual aid depicting the appearance of each cassette model to be maintained for propack packaging Another duplicate of this visual aid shall be maintained at the QC area facilitating accurate quality control measures | 2024-12-20 | Julie Bartoli
3 | Revise RCH-WI-000362 to from rev 1 to rev 2 to include a requirement for verifying the cassette being inspected against the visual aid in the QC area during QC inspection. | 2024-12-20 | Julie Bartoli
4 | Revise RCH-WI-000161 to from rev 1 to rev 2 to include a requirement for verifying the cassette being inspected against the visual aid in the QC area during QC inspection. | 2024-12-20 | Julie Bartoli

## Sheet2 (hidden)
CAPA# |  | Initiation Date:
Problem Statement | Owner:
|  | Description | Due Date
| Investigation: | <Root cause populated after completion>
| Action Planning: | N/A
| Action 1: | <Brief action description> - Owner name
| Action 2: | <Brief action description> - Owner name
| Action 3: | <Brief action description> - Owner name
| VOE 1: | <Brief VOE Description> - Owner name
| VOE 2: | <Brief VOE Description> - Owner name
| VOE 3: | <Brief VOE Description> - Owner name

## Sheet (Front)
CAPA-2024-000029 | Initiation Date: | 2024-08-27
Problem Statement | Owner: | Julie Bartoli
A recall is being initiated due to an error in which cases of PM Biopsy Cassettes (38441402) were erroneously packaged with PM Routine Cassettes (3 8441302). A recent complaint (CASE-2024-028945) received on 08/07/2024 alleged that the customer ordered PM Biopsy Cassettes - Green, (384414 02) and received a product that had the correct labeling but contained the wrong product - PM Routine Cassettes - Green (38441302) instead of PM Biopsy Cassettes - Green, (38441402) | Due Date
| Investigation | 2024-09-11
| Planning: | 2024-10-22
| VOE 1: | 2024-01-08
| VOE 2: | 2025-01-08